## Slide 1: Internal Quality Control Monitoring
Merih Tesfazghi, PhD, DABCC
Laboratory Director and Assistant Professor of Pathology, Rush Univeristy, Chicago IL
Laboratory Director, Johns Hopkins Clinical Research Project, Blantyre, Malawi
Co-Director, External Quality Assessment Program, Pathologists Overseas

## Slide 2: Why should labs care about quality?
Laboratory is a service provider entity.
With service delivery comes customer.
Customer satisfaction defines the future of labs.
The question then is how “well” should labs serve.

## Slide 3: What`s quality?
Philip Crosby
Four Absolutes of Quality Management
1979
Quality is meeting customer needs
Customer satisfaction is the minimum definition of quality
Notes: Perhaps a simple way to define quality is meeting or exceeding customer needs. To meet or exceed customer need, the lab must characterize the customer`s need. Before we understand their needs, we need to first know who our customers are. Know your customers, understand their need, meet or exceed need. We are going to discuss this as we move forward but as we all know labs have a range of customers with perhaps with a range of needs.

## Slide 4: Laboratory customers?
Community
Patients
Public Health
Laboratory
Physicians Health care
provider

## Slide 5
Labs produce “numbers”
Quality makes those numbers “interpretable”
Interpretable results have clinical relevance, which leads to trust
Glu (Scenarios) | #s | Quality check | Reported results
Scenario 1 | 47 | Glycolysis | -
Scenario 2 | 99 | None | 99
Scenario 3 | 600 | Contamination | -
Scenario 4 | 147 | QC out | -
Scenario 5 | 990 | Entry error | -
Why should clinicians trust lab results?
Glu (Scenarios) | #s | Quality check | Reported results
Scenario 1 | 47 | Glycolysis | 47
Scenario 2 | 99 | None | 99
Scenario 3 | 600 | Contamination | 600
Scenario 4 | 147 | QC out | 147
Scenario 5 | 990 | Entry error | 990
Lab 1
Lab 2

## Slide 6: Quality Management System (QMS)
Coordinated approach to direct and control activities of the laboratory
Everything a lab does to meet or exceed the needs and requirements of its clients
Quality control is a component of quality assessment in QMS, primarily designed to monitor analytical performance
Organization
Personnel
Equipment
Purchasing & Inventory
Process Control
Information Management
Documents & Records
Occurrence Management
Assessment
Process Improvement
Customer Service
Facilities & Safety

## Slide 7: Quality control (QC)
Surrogate samples that are intended to simulate clinical samples from patients.
Monitors a measurement procedure to verify that a method continues to meet performance specifications.
QC monitors the analytical performance of an assay.
  Does not monitor preanalytical testing phase
  Does not monitor postanalytical testing phase

## Slide 8
What do we need to consider when establishing a quality control system in our laboratories?

## Slide 9: Mitigating risk of harm to patients
Probability: How likely error occurs?
Severity: Degree of harm if error left undetected?
Detectability: How robust is the QC plan to detect error?
The primary goal of QC system is to mitigate the risk of harm to patients associated with reporting erroneous results.

## Slide 10: What does effective QC system entail?
1
QC material selection
2
QC target value and range
3
QC monitoring
4
QC frequency of analysis

## Slide 11: QC material selection
QC materials should be commutable.
QC concentrations should span AMR and clinical decision limits.
QC materials should be homogenous and stable.
QC material should have minimal vial-to-vial variabilities.
QC materials should have reasonable cost.

## Slide 12: Commercial QC materials
QC materials made and supplied by the manufacturers
  Optimized QCs for specific reagents or instruments
  Not optimal to detect systemic error if formulation mimics calibrator
QC materials made by a third party for the manufacturers
  Made based on specifications provided by the manufacturers
  May not detect systemic error if formulation mimics calibrator
QC materials made by a third party independent of manufacturers
  Made independently, not specific to instrument, reagent or calibrator
  QC materials different from calibrators

## Slide 13: Patient specimen QC samples
Specimen pool
  Challenging to cover AMR and clinical decision limits
  May require supplementing with purified materials
  Pooling and supplementing may introduce matrix bias
  Issue of stability for some measurands

## Slide 14: Forms of QC materials
Lyophilized QC
  Advantage: stable, longer shelf life, good for labs without freezers
  Challenge: requires reconstitution – introduces variability
Liquid QC
  Advantage: convenient – ready to use
  Challenge: requires freezer
Assayed QC
  Target and limit values provided by manufacturers
Unassayed QC
  Target and limit values NOT provided by manufacturers

## Slide 15
Global Survey: Types of QC materials commonly used?
(Global Survey by Westgard.com)

## Slide 16: Effective QC monitoring
1
QC material selection
2
QC target value and range
3
QC monitoring
4
QC frequency of analysis

## Slide 17: Appropriate QC target and acceptable range
Target value is an estimated mean concentrations of the QC obtained from a stable measurement system.
Range is a decision limit around the target value obtained from a stable measurement system.
2SD is commonly used to define ranges

## Slide 18: Establishing QC target and range values
Labs should establish their own target values and ranges
  Ideal target/range: 6-12 months data
  Initial range: 20 days
  Initial target: 10 days
  Provisional target: < 10 days
Old QC lot ranges can also be used and monitored closely
Manufacturer provided target values and ranges can be temporarily used.

## Slide 19
Global Survey: How do labs set QC limits?
(Global Survey by Westgard.com)

## Slide 20
Assume that following verification studies your laboratory began testing sodium on a new analyzer on Day 13. Immediately after this transition, clinicians noticed an increase in low sodium levels in otherwise clinically asymptomatic patients. A QC review of both levels did not reveal any apparent bias; however, further audit (pt results old vs new analyzers) confirmed an increase in low sodium with a median negative bias of 5 mmol/L. The root cause for the observed negative bias was identified and resolved.
Scenario
4 admissions
Intervention interruptions
Care rescheduling
Patient experience
Cost

## Slide 21: Effective QC monitoring
1
QC material selection
2
QC target value and range
3
QC monitoring
4
QC frequency of analysis

## Slide 22: Frequency of QC analysis
In the US (CLIA regulation 493.126(d)(3), at least once each day.
  QC levels
    Quantitative: 2 levels
    Qualitative: neg + pos
Major factors that determine QC frequency
  Risk of reporting error
  Time between QCs
  Test volume between QCs
Decisions are made by the Lab Director

## Slide 23: Frequency of QC analysis (Example)
https://www.qcnet.com/resources/qc-articles/the-frequency-of-quality-control-testing

## Slide 24: Effective QC monitoring
1
QC material selection
2
QC target value and range
3
QC monitoring
4
QC frequency of analysis

## Slide 25: Common QC monitoring Tools
Levey-Jennings (LJ) Chart
Westgard Rules

## Slide 26: LJ chart
A repetitive analysis of QC samples produce a Gaussian distribution.
68% of results will be within + 1 s of the mean.
95.5% of results will be within + 2 s of the mean.
99.7% of results will be within + 3 s of the mean.

## Slide 27: Manual preparation of LJ charts
https://www.westgard.com/lesson12.htm#4

## Slide 28: Westgard rules notation
NS
  N= # observation that exceeds “S” from X
  S= standard deviation
RS
  Range b/n two observations
  S= standard deviation
Levey-Jennings (LJ) Chart
Westgard Rules

## Slide 29: What are the outcomes of QC?
Was there error? | Was QC out? | Outcome
No | No | True Accept
No | Yes | False Reject
Yes | No | False Accept
Yes | Yes | True Reject
QC is either accepted or rejected
4 possible QC outcomes based on whether significant error was present and whether QC was “in” or “out”

## Slide 30
QC review rules are tradeoff b/n true and false rejection.
  Minimize false alerts
  Maximize true error detection
  Minimize harm to patients
True Error
False Error
Tradeoff between false and true alerts

## Slide 31: Westgard rules (examples)
Rules | Definition | Applied | Meaning
12s | 1 QC result exceeds ± 2 SD | Within run | Warning
22s | 2 consecutive QC results exceed ± 2SD | Within/across run | Reject
13s | 1 QC result exceeds ±3SD | Within run | Reject
R4s | 2 QC results span 4SD | Within run | Reject
41s | 4 results exceed 1 SD in the same direction | Within run/across | Reject
10x | 10 consecutive QC results fall on one side of the mean | Within run | Reject

## Slide 32: Shifts and Trends
Trend: gradual up-trending or down-trending of QC values.
Shift: sudden change of QC values from one side to another.
Both are caused by systematic error, but you may not always have to reject a run.
However, it`s critical to closely watch and investigate.

## Slide 33: Six Sigma metric
Six Sigma = TAE%(x) - |Bias%| (x)/ CV%(x)
  TAE: maximum error allowed without invalidating the results or adversely affecting medical decision-making.
  Bias: from a reference method. Usually imputed as zero.
  CV: precision from a replication experiment or routine QC data
  X: QC concentrations (level 1, 2, 3, etc)

## Slide 34
Six Sigma metric
Based on six sigma calculations, labs can decide on
  QC rules
  QC levels
  QC runs
6 sigma (1 rule): 13s, N=2 R=1
5 sigma (3 rules): 13s/22s/R4s N=2, R=1
4 sigma (4 rules): 13s/22s/R4s/41s (N=4, R=1 or N=2, R=2)
<4 sigma (5 rules): 13s/22s/R4s/41s/8x (N=4, R=2 or N=2, R=4)
6 sigma
Single rule
<4 sigma
Multiple rule
https://www.westgard.com/westgard-sigma-rules.htm

## Slide 35: QC review in the US
Frequency | Reviewer | Points of Evaluation
Daily | Technologist/Operator | QC rules Corrective action
Weekly | Senior Technologist | Mean and SD Shifts/Trends
Monthly | Senior Technologist and/or Manager and Director | Mean and SD Shifts/Trends Peer data (if available)

## Slide 36
Stop reporting patient results
Investigate and troubleshoot
Verify that patient results were not compromised
Address patient results that were compromised
How to respond to an out-of-control QC

## Slide 37: Investigation and troubleshooting: repeat
QC is Rejected
Repeat QC
Rejected again
Repeat QC
Rejected again
Repeat QC
Investigate
Identify type of errors
Systematic Error
Random Error
Start with simple and easy to solve potential causes
Right reagents?
New Calibration?
Work through the difficult potential causes
Instrument maintenance
Modification from last troubleshooting
New lot of QC?
Expired reagents?
Any changes since last acceptable QC?
Degradation of reagents

## Slide 38: How to address impacted patient results
Criteria for correction – determine clinical significance of impact
    Total allowable error
    Medical requirements
    Professional organizations/accrediting agencies
In the US, result correction entails calling/informing clinicians about the change
If result patient retesting is not possible because of stability (or other) issues, result correction still needs to occur
  Report results as invalid and inform clinicians

## Slide 39
Test every patient until last good QC
Test every patient until point of error
Test patients in batches until last good QC
Test patients in batches until point of error
Test patients in intervals until last good QC
Test patients in intervals until point of error
Ensure that system is now in control or retest patient specimens on a different instrument
How to identify extent of impact on patient results?

## Slide 40: QC limitations
Noncommutability
  QC materials may not behave like real patient specimens
  Lowers sensitivity and specificity of error detection
A snapshot of moment
  Assumes method remains stable between QC events
  Does not monitor performance in real-time
Expensive
  Large proportion of total testing cost
Does not challenge the entire testing process

## Slide 41: QC limitation: snapshot of a moment
QC
QC
1st Shift
2nd Shift
3rd Shift
L1/L2
L1/L2
L1/L2
QC
QC results
Patient results
QC
QC
QC
QC
QC
QC

## Slide 42: Patient data as QC check - moving average
QC
QC
QC
QC
QC
QC
QC results
Patient results
Moving average

## Slide 43: Patient data as QC check - Delta
Difference between current and previous results within a given time.
Delta triggers investigation before patient result is reported.
Delta must balance false and true alerts.
Unlike QC, delta can detect preanalytical and postanalytical errors.

## Slide 44: Questions?
Thank you